## Sheet1
Your Company Name and/or Logo |  |  |  |  | Your Address
CALIBRATION REPORT
Mfr.: | Rosemount |  |  | Cal Date: | 2013-05-24
Model No.: | 3051CD |  |  | Cal Tech: | My Name
Serial No.: | 123456789 |  |  | Std. No. 1: | AJ02671
Tag No.: | LT101 |  |  | Std. No. 2: | 176608
Location: | Plant #1 |  |  | Std. No. 3:
Description: | Level Xmtr |  |  | Procedure: | SOP-1
Range/Units: | 250 | "H2O |  | Revision: | A
Input LRV: | 0 |  |  | Rev. Date:
Input URV: | 250 |  |  | Process Tol.: | 1%
Output/Units | 4-20 | mA |  | Note: | T.A.R. < 4:1
Output LRV: | 4 |  |  | Cal Interval: | 12 | Months
Output URV: | 20 |  |  | Cal Due Date: | 2014-05-24
AS-FOUND DATA
Standard | Instrument | Low Limit | High limit | Instrument | Deviation | Error
"H2O | mA | mA | mA | "H2O | "H2O | % of Span
0 | 4.1 | 3.84 | 4.16 | 1.5625 | 1.5625 | 0.625
50 | 7.32 | 7.04 | 7.36 | 51.875 | 1.875 | 0.75
100 | 10.55 | 10.24 | 10.56 | 102.3438 | 2.3438 | 0.9375
150 | 13.79 | 13.44 | 13.76 | 152.9688 | 2.9688 | 1.1875 | F
200 | 17.05 | 16.64 | 16.96 | 203.9062 | 3.9062 | 1.5625 | F
250 | 20.3 | 19.84 | 20.16 | 254.6875 | 4.6875 | 1.875 | F
200 | 17.06 | 16.64 | 16.96 | 204.0625 | 4.0625 | 1.625 | F
150 | 13.8 | 13.44 | 13.76 | 153.125 | 3.125 | 1.25 | F
100 | 10.56 | 10.24 | 10.56 | 102.5 | 2.5 | 1 | F
50 | 7.33 | 7.04 | 7.36 | 52.0312 | 2.0312 | 0.8125
0 | 4.1 | 3.84 | 4.16 | 1.5625 | 1.5625 | 0.625
F in right margin indicates Failed Cal (Out of Tol.)
AS-LEFT DATA
Standard | Instrument | Low Limit | High Limit | Instrument | Deviation | Error
"H2O | mA | mA | mA | "H2O | "H2O | % of Span
0 | 4 | 3.84 | 4 | 0 | 0 | 0
50 | 7.19 | 7.04 | 7.36 | 49.8438 | -0.15625 | -0.0625
100 | 10.4 | 10.24 | 10.56 | 100 | 0 | 0
150 | 13.59 | 13.44 | 13.76 | 149.8438 | -0.15625 | -0.0625
200 | 16.8 | 16.64 | 16.96 | 200 | 0 | 0
250 | 19.99 | 19.84 | 20.16 | 249.8437 | -0.15625 | -0.0625
200 | 16.81 | 16.64 | 16.96 | 200.1562 | 0.15625 | 0.0625
150 | 13.6 | 13.44 | 13.76 | 150 | 0 | 0
100 | 10.41 | 10.24 | 10.56 | 100.1562 | 0.15625 | 0.0625
50 | 7.2 | 7.04 | 7.36 | 50 | 0 | 0
0 | 4.01 | 3.84 | 4.16 | 0.15625 | 0.15625 | 0.0625
Cal Notes: | Out of tol as found
Cal to process tol as left
Serial No.: | 123456789
Calibrated by: | My Name |  |  | Approved by: | My Super
Cal Date: | 2013-05-24 |  |  | Date: | 2013-05-24
Formulas:
My Name = EDATE(F6,[My Name, row above])
DEMO